**FILT’RAY in-line filter, 4 months**

Non-sterile

Reference: **AT030744**

Specification description

AQUATOOLS anti-Legionella and all-germ in-line filter.

Suitable for filtering water to supply equipment used in the treatment of medical devices (e.g. endoscopes).

Device made from MBS.

Membrane type: tubular microfiltration.

Membrane made from polyethylene.

Pore cut-off threshold: 0.1μm nominal and 0.2μm absolute sterilising grade.

Qualified water volume per duration of use: 6,000 litres.

Filtration flow rate: 6.3 lpm at 3 bar and 2.2 lpm at 1 bar (filter only, excludes any flow rate restriction in the mixer/tap).

Maximum upstream pressure at the point-of-use: 5 bar.

Maximum lifespan: 4 months after installation.

Resistance to thermal shocks: 70°C for 30 minutes.

Resistance to chemical shocks: active chlorine levels up to 100 ppm for 4 hours (cumulative total).

Non-sterile filter.

Medical device CE Class I.

Packaged in boxes of 10 pieces, each filter individually wrapped.